## cable-1
CABLE-1 (without blue label)
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 7bis | 8 | 8bis | 9 | 9bis | 10 | 11 | 11bs | 12 | 12bis | 13 | 14 | 15 | 16 | 17 | 17bis | 18 | 18bis | 19 | 20
0 | [27 empty] | 0 | bar: CABLE-1 (without blue label) 9; CABLE-1 (without blue label) 11; CABLE-1 (without blue label) 12
1000 | 89 | 89 | 89 | 88 | 89 | 89 | 89 | 88 | 90 | 89 | 90 |  | 90 | 88.8 | 89 | 89.2 | 88 | 89.1 | 88.6 | 89 | 88.9 |  | 89 | 90.1 | 89 | 89.8 | 89.9 | 1000
5000 | 448 | 450 | 452 | 448 | 451 | 450 | 452 | 446 | 458 | 449 | 454 | 450 | 454 | 447.3 | 451 | 449.8 | 447 | 451.1 | 446.9 | 450.4 | 448.5 | 326.5 | 450 | 455.6 | 448 | 453.6 | 453.9 | 5000
7000 | 630 | 633 | 636 | 630 | 635 | 632 | 636 | 628 | 645 | 632 | 639 | 634 | 639 | 629 | 630 | 633 | 635 | 635 | 629 | 632 | 631 | 535 | 632 | 642 | 635 | 638 | 639 | 7000
8000 | 724 | 727 | 731 | 724 | 729 | 726 | 730 | 722 | 742 | 726 | 732 | 729 | 732 | 722 | 725 | 726 | 729 | 729 | 722 | 727 | 724 | 641 | 725 | 738 | 728 | 731 | 733 | 8000
9000 | 814 | 818 | 824 | 816 | 823 | 819 | 819 | 816 |  | 820 | 825 | 823 | 826 | 816 | 818 | 820 | 823 | 824 | 816 | 822 | 818 | 724 | 818 | 835 | 822 | 827 | 828 | 9000
9500 | 864 | 868 | 872 | 864 | 870 | 866 | 866 | 865 |  | 869 | 874 | 873 | 875 | 863 | 868 | 868 | 872 | 874 | 866 | 873 | 869 | 789 | 867 | 892 | 871 | 878 | 878 | 9500
10000 | 912 | 916 | 921 | 912 | 917 | 913 | 912 | 912 |  | 917 | 923 | 921 | 924 | 912 | 915 | 917 | 920 | 922 | 913 | 920 | 916 | 822 | 914 | 943 | 918 | 925 | 926 | 10000
10500 | 960 | 964 | 969 | 960 | 964 | 960 | 960 | 960 |  | 965 | 971 | 969 | 972 | 961 | 964 | 967 | 969 | 971 | 962 | 969 | 964 | 887 | 962 | 995 | 967 | 974 | 975 | 10500
960
[7 empty] | 8.004 |  | 8.005 | [12 empty] | 8 |  | 8.0001
[22 empty] | 0 |  | 89.2
[21 empty] | fluctuates | 217 |  | 448
[9 empty] | CABLE-1-4-2 | [12 empty] | CABLE-1 - 9-
Note on H18: Ch triped at 15 KV without load
Note on J18: Trips at 15KV (without load)
Note on L14: triped at 15KV without load
Note on O12: tripped at 15KV without load
Note on Q12: tripped at 15KV without load
Note on W18: ZAGANIDIS: at 15KV without loasd: 1-9-1 1-9-2

## cable-2
CABLE 2 (with blue label)
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
0 | [20 empty] | 0
1000 | 89 | 89 | 88 | 89 | 88 | 89 | 90 | 89 | 89 | 90 | 89 | 89 | 89 | 88 | 89 | 89 | 89 | 90 | 90 | 90 | 1000
5000 | 446 | 449 | 451 | 447 | 450 | 448 | 450 | 458 | 453 | 453 | 448 | 451 | 452 | 448 | 451 | 450 | 452 | 460 | 452 | 455 | 5000
7000 | 628 | 632 | 635 | 629 | 634 | 631 | 634 | 645 | 638 | 638 | 630 | 633 | 635 | 630 | 635 | 632 | 635 | 647 | 639 | 639 | 7000
8000 | 721 | 725 | 728 | 721 | 726 | 723 | 727 | 740 | 731 | 731 | 722 | 726 | 729 | 722 | 727 | 724 | 728 | 742 | 732 | 733 | 8000
9000 | Fluct. | fluct. | 824 | 816 | 822 | 819 | 823 | 840 | 825 | 826 | 815 | 819 | 823 | 815 | 821 | 818 | 822 | 839 | 826 | 827 | 9000
9500 |  |  | 871 | 863 | 870 | 866 | 870 | 890 | 874 | 875 | 863 | 867 | 871 | 863 | 870 | 866 | 870 | 889 | 875 | 876 | 9500
10000 |  |  | 919 | 911 | 918 | 914 | 919 | 941 | 923 | 924 | 911 | 915 | 920 | 911 | 918 | 914 | 918 | 939 | 923 | 924 | 10000
10500 |  |  | 967 | 958 | 966 | 962 | 966 | 990 | 971 | 972 | 962 | 966 | 971 | 962 | 969 | 964 | 969 | 995 | 974 | 975 | 10500
| 4.004 | 4.005 | 8 | 8.001 | 8.002 | 8.003 | 8.004 | 8.005 | 12 | 12.001
Cable-2-1
Note on B21: 15KV without load: 2-1-1 trips 2-1-2

## ISR HV cables
HV-xx : Tripolar male - jupiter B & T
short-xx : tripolar femelle -jupiter B&T : L~1.5m
| HV-13 + Short cable 3 (tested:OK) |  | HV-07 + Short cable 1 (tested :OK) |  | HV-05+ Short cable 4 (tested: OK) |  | HV-02 + Short cable 3(tested :OK) |  | HV-04 + Short cable 4 (tested :OK) |  | HV-06 + Short cable 3(tested :OK) |  | HV-07 + Short cable 1 (tested :OK) |  | HV-09 + Short cable 1 (tested :OK) |  | HV-12 + Short cable 3 (tested :OK) |  | HV-3 |  | HV-11 |  | short cable 3 |  | short cable 1 |  | short cable 4 |  | short cable 2
| B | T | B | T | B | T | B | T | B | T | B | T | B | T | B | T | B | T | B | T | B | T | B | T | B | T | B | T | B | T
100 | 9.7 | 8.9 | 9.5 | 9.2 | 9.5 | 11 | 9.5 | 9.4 | 9.6 | 8.8 | 9.5 | 9.3 | 9.5 | 11 | 8.8 | 9 | 8.8 | 8.6 | TESTED 20 oct 2011: OK but data lost |  | TESTED 20 oct 2011: OK but data lost
1000 | 88.8 | 89.1 | 89.7 | 88.9 | 89.4 | 89.3 | 89.3 | 89.5 | 88.8 | 89.1 | 89.7 | 89 | 89.4 | 89.3 | 89.6 | 89.6 | 89.1 | 89.5
5000 | 447.9 | 449.9 | 452.7 | 449.3 | 451.6 | 450.3 | 448 | 450 | 447 | 449.1 | 451.9 | 448.3 | 450.7 | 449.8 | 451 | 450 | 448.2 | 450.5
7000 | 631.1 | 634.1 | 637.9 | 633.5 | 636.3 | 634.2 | 631 | 634 | 632.9 | 635.3 | 639.2 | 634.1 | 636.8 | 634.8 | 636.1 | 634.7 | 631.1 | 634.1
8000 | 724.6 | 727.7 | 732.3 | 727 | 730.2 | 727.6 | 726 | 729 | 724.7 | 727.6 | 732.5 | 726.6 | 729.9 | 727.5 | 728.5 | 727 | 723.7 | 726.8
9000 | 819.7 | 822.8 | 828.5 | 822.2 | 825.5 | 822.4 | 819 | 822 | 818.5 | 821.7 | 827.7 | 821.1 | spark | spark | 825.1 | 822.6 | 816.8 | 820.8
9500 | 869.2 | 872.1 | 878.2 | 871.2 | 874.5 | 870.9 | 866 | 869 | 867 | 870 | 876.3 | 869.4 |  |  | 873.1 | 870.2 | 865.8 | 869.1
10000 | 917.9 | 920.8 | 927.5 | 920.1 | 923.4 | 919.3 | 914 | 917 | 915 | 918 | 925 | 917.6 |  |  | 921.4 | 918 | 913 | 916.4
10500 | 967.1 | 970.1 | 977.6 | 970 | 973.5 | 968.8 | 961 | 964 | 964 | 967.5 | 974.7 | 966.8 |  |  | 970.6 | 966.7 | 962 | 965
[9 empty] | ch 1, 2 | [5 empty] | ch 4,5 |  | ch 2,3
Note on E5: 15KV : discharge without short cable
Note on L6: 15KV test: Big descharge at 15KV
Note on Q4: load OK 15KV sparks ( 1psark/30 min)

## HV modules
cables : | Cable-2-7-1 | Cable-2-7-2 | Cable-2-8-1 | Cable-2-8-2 | Cable-2-9-1 | Cable-2-9-2 | cables : | Cable-2-7-1 | Cable-2-7-2 | Cable-2-8-1 | Cable-2-8-2 | Cable-2-9-1 | Cable-2-9-2
Module 1910029 | [6 empty] | Module 1910010
| ch 0 | ch 1 | ch 2 | ch 3 | ch 4 | ch 5 |  | ch 0 | ch 1 | ch 2 | ch 3 | ch 4 | ch 5
0 | [6 empty] | 0
1000 | 88 | 89 | 89 | 88 | 89 | 89 | 1000 | 88 | 89 | 89 | 88 | 89 | 88
5000 | 446 | 448 | 450 | 447 | 450 | 448 | 5000 | 448 | 449 | 451 | 447 | 450 | 448
9000 | 809 | 814 | 817 | 812 | 818 | 814 | 9000 | 813 | 814 | 819 | 812 | 817 | 814
10500 | 950 | 956 | 960 | 954 | 961 | 956 | 10500 | 954 | 955 | 961 | 954 | 959 | 956
cables : | Cable-2-7-1 | Cable-2-7-2 | Cable-2-8-1 | Cable-2-8-2 | Cable-2-9-1 | Cable-2-9-2 | cables : | Cable-2-7-1 | Cable-2-7-2 | Cable-2-8-1 | Cable-2-8-2 | Cable-2-9-1 | Cable-2-9-2
Module 1910023 | [6 empty] | Module 1910008
| ch 0 | ch 1 | ch 2 | ch 3 | ch 4 | ch 5 |  | ch 0 | ch 1 | ch 2 | ch 3 | ch 4 | ch 5
0 | [6 empty] | 0
1000 | 88 | 89 | 89 | 88 | 89 | 89 | 1000 | 88 | 89 | 89 | 88 | 89 | 89
5000 | 446 | 448 | 449 | 446 | 450 | 448 | 5000 | 448 | 449 | 450 | 447 | 450 | 450
9000 | 811 | 815 | 818 | 813 | 818 | 815 | 9000 | 811 | 814 | 817 | 811 | 816 | 816
10500 | 952 | 956 | 961 | 954 | 960 | 957 | 10500 | 956 | 959 | 962 | 956 | 961 | 961
cables : | Cable-2-7-1 | Cable-2-7-2 | Cable-2-8-1 | Cable-2-8-2 | Cable-2-9-1 | Cable-2-9-2
Module 1910015
| ch 0 | ch 1 | ch 2 | ch 3 | ch 4 | ch 5
0
1000 | 88 | 89 | 89 | 88 | 89 | 89
5000 | 446 | 448 | 449 | 447 | 450 | 448
9000 | 812 | 815 | 816 | 812 | 820 | 811
10500 | 952 | 956 | 958 | 953 | 960 | 953

## Gaps
| cm | cm | mb
Kodel-cms-RE4/3-BO2
1cm =1 mb
Leak test at 20mb | column1 | column 2 | P - Patm | Time | DP (mb)
| 12.8 | 33.2 | 20.4 | 11:20
| 13.8 | 32.1 | 18.3 | 11:30 | 2.1
| 14.1 | 31.2 | 17.1 | 11:40 | 1.8
| 15.5 | 30.5 | 15 | 11:50 | 1.5
| 16.1 | 29.9 | 13.8 | 12:00 | 1.2
Spacers | 170 |  | for all spacers : 0,0
For middle
Kodel-cms-RE4/3-TWO2
while leak test at 20mb cracking sound was heard
Kodel-cms-RE4/3-TWO3
Leak test at 10mb | column1 | column 2 | P - Patm | Time | DP
| 17.6 | 28.1 | 10.5 | 5:40
| 20.5 | 25.3 | 4.8 | 5:50 | 5.7
| 22.5 | 23.3 | 0.8 | 5:57 | 4
| 23 | 23 | 0 | 6:00 | 0.8
spacers | 45 | for all spacers : 0,0
| 5 | for all middle all 1,1
Kodel-cms-RE4/3-TN02
Leak test at 10mb | column1 | column 2 | P - Patm | Time | DP
| 18.2 | 27.6 | 9.4 | 6:51
| 20 | 25.9 | 5.9 | 7:01 | 3.5
| 21.4 | 24.4 | 3 | 7:11 | 2.9
spacers | 114 | ® | 114 x (0,0)
middle | 10 | ® | 10 x (1,1)
Kodel-cms-RE4/3-TW01
Leak test not done
spacers | 45 | ® | 44 x(0,0) + 1 x(1,1)
Middle | 8 | ® | 8 x (1,1)
Declared as BAD from KOREA:
KODEL-CMS-RE4-R3-TN01
KODEL-CMS-RE4-R3-B03
KODEL-CMS-RE4-R3-TW03

## HV_scan_Kodel_gaps
Gas | 95.5 freon | 4.2% Iso | 0.3 SF6
site : | ISR
Date : | 25/10/2010
A1526N-1910019
[5 empty] | Ch0 | ch1 | Ch2
[5 empty] | HV3-shortcable_2_B | HV3-shortcable_2_T | HV11-shortcable_1_B | HV11-shortcable_1_B
| HV0SET | Hvmon |  | Time | RE4_R3_B02 | RE4_R3_TW03 | RE4_R3_TN02 | RE4_R3_TN02
| 0.1 | 111 |  | 10 | 0 | 0 | 0.01 | 0.01 / 107 V | 0915 (23 Oct., 2010)
| 1 | 999 |  | 120 | 0.01 | 0.03 | 0.06 | 0.07 / 1000 V | 9:17
| 2 | 2001 |  | 120 | 0.03 | 0.06 | 0.13 | 0.13 / 2000 V | 9:38
| 3 | 3001 |  | 240 | 0.05 | 0.1 | 0.19 | 0.19 / 3001 V | 9:44
| 4 | 4000 | 20:15 | 240 | 0.08 ® 0.06 (at 22.26 h) | 0.14 | 0.25 | 0.25 / 4002 V | 9:50
| 5 | 5001 | 22:30 | 600 | 0.09 | 0.19 | 0.32 | 0.32 / 5001 V | 9:55
| 6 | 6001 | 22:42 | 600 | 0.1 | 0.25 | 0.42 | 0.38 / 6002 V | 10:10
| 7 | 7000 | 22:57 | 600 | 0.12 | 0.35 | 0.76 (TRIPPED) | 0.44 / 7001 V | 10:25
| 8 | 8001 | 23:10 | 600 | 0.16 | 0.59 |  | 0.48 / 7500 V | 10:40
| 8.5 | 8501 | 23:23 | 900 | 0.2 | 0.88 | ramped again and tripped | 0.52 / 8001 V | 10:50
| 9 | 9001 | 23:40 | 900 | 0.26 | 1.41 | again at 7 kV | 0.57 / 8501 V | 11:00
| 9.1 | 9101 | 23:57 | 900 | 0.27 | 1.55 |  | 0.63 / 9000 V | 11:10
| 9.2 | 9201 | 0:14 | 900 | 0.29 | 1.68 |  | 0.64 / 9101 V | 11:20
| 9.3 | 9300 | 0:31 | 900 | 0.31 | 1.86 | No gas flow at all in ch # 7 | 0.65 / 9201 V | 11:30
| 9.4 | 9400 |  |  | 0.31 ® 0.27 at 9:30 | 2.04 | even after the gas flow | 0.66 / 9300 V | 11:40
[5 empty] | on 22 Oct. 2010 |  | was rectified by Roberto | 0.68 / 9400 V | 11:50 | ® 0.66 / 9400 V
[7 empty] | at 1900 h on 22 Oct. 2010 |  |  | at 10:08 h
on 24 oct. 2010
gas channel # 5 connected
to narrow gap which was
earlier connected
to bottom gap
at 2105 h (22 Oct 2010)
flow kept at 2 lph

## HV_scan_GT_gaps
GAS : | [7 empty] | GAS :
Mixture: | 95.5 freon | 4.2% Iso | 0.3 SF6 |  |  |  |  | Mixture: | 95.5 freon | 4.2% Iso | 0.3 SF6
Gas supplied since : | 24/10/2010 | [6 empty] | Gas supplied since :
GAS-CH10 |  | 4 l/h | bubbler gas leak test : OK |  |  |  |  | GAS-CH1 |  | 4 l/h | bubbler gas leak test : OK
GAS-CH3 |  | 4 l/h | bubbler gas leak test : OK |  |  |  |  | GAS-CH5 |  | 4 l/h | bubbler gas leak test : OK
GAS-CH4 |  | 4 l/h | bubbler gas leak test : OK |  |  |  |  | GAS-CH |  | 4 l/h | bubbler gas leak test : OK
GAPS : | [7 empty] | GAPS :
Bakelite: | [7 empty] | Bakelite:
Supplier : | GT | [6 empty] | Supplier : | GT
GT-344 | S.N.: 344 | Type: R3-Bottom | Supplier : GT | [5 empty] | S.N.: 347 | Rectangular | Supplier : GT
GT-345 | S.N.: 345 | Type: R3-T-W | Supplier : GT | [5 empty] | S.N.: 348 | Rectangular | Supplier : GT
GT-342 | S.N.: 342 | Type: R3-T-N | Supplier : GT | [5 empty] | S.N.: 346 | Rectangular | Supplier : GT
ENV : | [7 empty] | ENV :
T | 23 | [6 empty] | T
H20 | 31% | [6 empty] | H20
P | 963.26 | [6 empty] | P
Site | ISR | [6 empty] | Site | ISR
|  | Date : 28/10/2010 | [7 empty] | Date :
|  | A1526N-1910019 | [7 empty] | A1526N-xxxxxx
|  | Ch0 |  | Ch1 |  | Ch2 |  |  |  | Ch |  | Ch |  | Ch
|  | HV3-shortcable_2_B |  | HV3-shortcable_2_T |  | HV11-shortcable_1_B
|  | GAS-CH10 |  | GAS-CH3 | [5 empty] | GAS-CH1 |  | GAS-CH5 |  | GAS-CH
|  | GT-344 |  | GT-345 |  | GT-342 |  |  |  | GT-347 |  | GT-348 |  | GT-346
Time (min) | HV0SET | Hvmon | Imon | Hvmon | Imon | Hvmon | Imon | Time (min) | HV0SET | Hvmon | Imon | Hvmon | Imon | Hvmon | Imon
1 | 0.1 | 111 | 0 | 98 | 0.01 | 107 | 0.01 | 1 | 0.1
5 | 1 | 998 | 0 | 996 | 0.02 | 1000 | 0.01 | 5 | 1
5 | 2 | 2000 | 0 | 1997 | 0.03 | 2000 | 0.01 | 5 | 2
5 | 3 | 3000 | 0 | 2997 | 0.04 | 3000 | 0.02 | 5 | 3
5 | 4 | 4000 | 0.01 | 3998 | 0.05 | 4001 | 0.02 | 5 | 4
5 | 5 | 5000 | 0.01 | 4997 | 0.06 | 5001 | 0.02 | 5 | 5
5 | 6 | 6000 | 0.02 | 5999 | 0.07 | 6002 | 0.02 | 5 | 6
10 | 7 | 7000 | 0.04 | 6998 | 0.08 | 7001 | 0.03 | 10 | 7
10 | 8 | 8001 | 0.06 | 7997 | 0.1 | 8001 | 0.05 | 10 | 8
10 | 8.5 | 8500 | 0.1 | 8498 | 0.11 | 8501 | 0.06 | 10 | 8.5
10 | 9 | 8999 | 0.18 | 8998 | 0.12 | 9000 | 0.1 | 10 | 9
15 | 9.1 | 9100 | 0.19 | 9097 | 0.13 | 9101 | 0.11 | 15 | 9.1
15 | 9.2 | 9201 | 0.21 | 9198 | 0.14 | 9201 | 0.12 | 15 | 9.2
15 | 9.3 | 9300 | 0.23 | 9298 | 0.14 | 9300 | 0.12 | 15 | 9.3
15 | 9.4 | 9400 | 0.25 | 9397 | 0.14 | 9400 | 0.14 | 15 | 9.4
15 | 9.5 | 9501 | 0.28 | 9497 | 0.14 | 9501 | 0.15 | 15 | 9.5
15 | 9.6 | 9600 | 0.31 | 9597 | 0.15 | 9601 | 0.17 | 15 | 9.6
15 | 9.7 |  |  | 9700 | 0.15 | 9701 | 0.19 | 15 | 9.7
| 9.8 |  |  | 9797 | 0.16 | 9801 | 0.21
offset ch0/1 :-0.02
offset ch1/2 :0.0
Note on D38: Imon fluctuates 0.18 -- 0.21
Note on D39: Imon fluctuates 0.20 -- 0.23
Note on D40: Imon fluctuates 0.22 -- 0.26
Note on D41: Imon fluctuates 0.25 -- 0.30
Note on D42: Imon fluctuates 0.27 -- 0.31
Note on D43: Imon fluctuates 0.30 -- 0.34 decharges?
Note on H44: Imon fluctuates time to time 0.17-0.22 decharges?
Note on H45: Imon fluctuates 0.19-0.22 decharges?

## NPD-BARC-11
SN : NPD-BARC-11
Gaps : | Kodel-cms-RE4/3-BO2 | Kodel-cms-RE4/3-TN02 | Kodel-cms-RE4/3-TWO3
Top | gas channel : 5
Bottom | gas channel :1
Gas flow | 4l/h
T | 22.4
H20 | 26%
P | 957.64 mb
Gas | 95.5 freon | 4.2% Iso | 0.3 SF6
site : | ISR
Date : | 25/10/2010
A1526N-1910019
[5 empty] | Ch0 |  | ch1 |  | Ch2
[5 empty] | HV3-shortcable_2_B |  | HV3-shortcable_2_T |  | HV11-shortcable_1_B
| HV0SET |  | Time | Hvmon | KODEL-CMS-RE4-R3-B02 | Hvmon | KODEL-CMS-RE4-R3-TN02 | Hvmon | KODEL-CMS-RE4-R3-TW03
| 0.1 |  | 120 | 111 | 0 | 99 | 0.01 | 107 | 0.01
| 1 |  | 120 | 999 | 0.02 | 996 | 0.06 | 1000 | 0.04
| 2 |  | 120 | 2000 | 0.04 | 1998 | 0.11 | 2000 | 0.07
| 3 |  | 120 | 3000 | 0.04 | 2997 | 0.16 | 3001 | 0.1
| 4 | 20:15 | 120 | 4000 | 0.07 | 3998 | 0.22 | 4001 | 0.13
| 5 | 22:30 | 120 | 5001 | 0.1 | 4998 | 0.28 | 5001 | 0.17
| 6 | 22:42 | 120 | 6000 | 0.12 | 5999 | 0.34 | 6002 | 0.22
| 7 | 22:57 | 120 | 7000 | 0.14 | 6998 | 0.4 | 7001 | 0.32
| 8 | 23:10 | 120 | 8001 | 0.17 | 7997 | 0.48 | 8001 | 0.56
| 8.5 | 23:23 | 120 | 8501 | 0.21 | 8498 | 0.52 | 8501 | 0.87
| 9 | 23:40 | 120 | 9001 | 0.23 | 8998 | 0.58 | 9001 | 1.42
|  |  |  | 9001 | 0.22 | 8998 | 0.55 | 9001 | 1.27
| 9.1 | 23:57 | 120 | 9101 | 0.23 | 9098 | 0.56 | 9102 | 1.39
| 9.2 | 0:14 | 300 | 9201 | 0.25 | 9198 | 0.57 | 9201 | 1.53
| 9.3 | 0:31 | 300 | 9300 | 0.26 | 9298 | 0.59 | 9301 | 1.67
| 9.4 |  | 300 | 9401 | 0.28 | 9397 | 0.6 | 9401 | 1.82
| 9.5 |  | 300 | 9501 | 0.3 | 9497 | 0.62 | 9502 | 1.95
| 9.6 |  | 300 | 9601 | 0.32 | 9598 | 0.63 | 9601 | 2.09
| 9.7 |  | 300 | 9701 | 0.34 | 9697 | 0.64 | 9701 | 2.26