## Sheet1
GAS :
Mixture: | 100% Ar | [7 empty] | Gap Area |  |  |  | Active Area
Gas supplied since : | 23/10/2010 at 5 a.m. | [6 empty] | L(cm) | l(cm) | h(cm) | S(m2) | L(cm) | l(cm) | h(cm) | S(m2)
GAS-5 |  | 5 l/h | GT-344 | S.N.: 344 | Type: R3-Bottom | Supplier : GT | Leaking in GIF (no bubbling in vertical position) after attempts to repair | 126.5 | 92.7 | 193 |  | 121.1 | 88.2 | 188 | 1.966
GAS-CH6 |  | 5 l/h | GT-345 | S.N.: 345 | Type: R3-T-W | Supplier : GT | bubbler gas leak test : OK | 126.2 | 117 | 52.6 |  | 120.8 | 112.5 | 47.6 | 0.555
No channel available |  | NA | GT-342 | S.N.: 342 | Type: R3-T-N | Supplier : GT | not connected | 116.9 | 93 | 136.4 |  | 111.4 | 88.4 | 131.4 | 1.313
| GT-344 | GT-345
HV cable | HV-2-1-2 | HV-2-1-1
| source : OFF |  | source : ON
| GT-344 | GT-345 | GT-344 | GT-345 ON
High Voltage | Current | Current | Current | Current
1500 | 0 | 0 | 0 | 0
1600 | 0 | 0 | 0 | 0
1700 | 0.2 | 0 | 0.2 | 0
1800 | 0.7 | 0 | 0.4 | 0
1850 | 1 | 0 | 0.8 | 0
1900 | 1.7 | 0.1 | 1.3 | 0.1
1950 | 3.1 | 0.6 | 2.3 | 0.6
2000 | 7.5 | 2 | 5.3 | 1.6
2050 | 14.5 | 3.7 | 10.3 | 2.7
2100 | 22 | 5 | 16 | 3.9
2150 | 29 | 6 | 21.7 | 5
2200 | 37 | 7 | 28 | 6
2250 | 45 | 9 | 34.7 | 7
2300 | 53 | 10 | 41.2 | 8.2
2350 | 61 | 11 | 48.6 | 9.2
2400 | 70 | 12 | 55.8 | 10.4
2450 | 78 | 13 | 63 | 11.4
2500 | 87 | 15 | 71 | 12.6
2550 | 96 | 16 | 78.5 | 13.7
2600 | 106 | 17 | 87 | 14.7
2650 | 114 | 18 | 95.5 | 15.7
2700 | 125 | 19 | 104.7 | 16.8
2750 | 134 | 21 | 114 | 17.9
2800 | 144 | 22 | 124 | 19
2850 | 156 | 23 | 134 | 20.3
2900 | 166 | 24 | 144.4 | 21.4
2950 | 178 | 25 | 155 | 22.4
R (ohm) | 5701254.2759 | 41152263.3745 | 6506180.8718 | 45871559.633
Ro (ohm x cm) | 280216647662.486 | 570987654320.988 | 319778789850.358 | 636467889908.257