Introduction for the Gas flow control

Title

Gas flow control using feedback on a precision needle valve. To remotely control gas flow, with one channel, for balanced distribution or , with numerous channels, to mix different gases to produce a required mixture.

Principle or aim

There are 2 basic approaches

Firstly, use a flow sensor feedback through microprocessor and servo motor to control the valve of a classic rotameter.

Secondly calibrate the valve with an accurate flow sensor, such as a MFC (Mass Flow Controller), and then correlate the number of turns on the valve to the flow rate.

The normal operation of a rotameter gives a not so precise analogue e reading requiring constant human intervention to remove drift in time. It cannot damp out the common cases of oscillation of the ball in the glass metered column (glass tube) giving another source of error.

Ther are many sources of flow measurement such as photo or inductive to read out the position on the rotameter. Similar devices used a few decades ago as an alarm for the required flow limits. Today LED strips associated to photodiodes/resistors could pick up the position of the xxx I the rotameter.

In addition, there are other fundamental ways such as that used in the MFC namely the heat extraction as a function of the flow rate of the gas.

Although relying on the glass graduated column errors are removed by calibration to a known accurate source eg; MFC and oscillations can be damped out.

In the pharmaceutical domain bubble counting, digitally, has been introduced. This would work for very small values of flow rate. A precise understand of the bubble volume would be required as a function of the bubbler geometry and liquid viscosity.

Faced with the need to make rapid progress we resorted to a high volume flow rate sensor, at least an order of magnitude over what would be required for a typical lab set up such as in the case of the RPCs , namely <10litres/hour. This was sourced through the gas group at CERN. This MEMS is a new technology widely available at a reasonable cost and as we will see, gives an amazing resolution in the order of 0.01 l/hr or better