## Sheet1
11.5
Muhammad, Ian, Zubayda
2023-12-13
12.39pm
NOVELEK 8174 |  | Set Point | Analog Reading | Monitor Reading | Analog Reading | Monitor Current |  | ECL | Notes
Serial No. 8901281 |  | [kV] | [kV] | [-kV] | [mA] | [V]
Type. HT8174 |  | 0 |  | 0.01 |  |  |  | 0.91
Module 34 [first] |  | 1 |  | 0.81 |  | 0.012 |  | 0.85
Channel B |  | 2 | 1.5 | 1.603 | 0 | 0.025 |  | 0.77
|  | 3 | [6 empty] | overload voltage at 2kV, load disconnected
Restarting test with load disconnected
|  | 1 |  | 0.806 |  | 0.01 |  | 0.85
|  | 2 |  | 1.603 |  | 0.024 |  | 0.85
|  | 3 |  | 2.341 |  |  |  |  | overload voltage, automatically ramping down
Module 2 |  | 0 | [8 empty] | Resistance
Channel A Test 1 |  | 1 |  | 0.784 |  | 0.832 |  | 0.98 | ECL change t 1.01 for current reading |  | 0.94231
|  | 2 |  | 1.583 |  | 1.571 |  | 1 |  |  | 1.0076
|  | 3 |  | 2.381 |  | 2.348 |  | 1.04 | ECL change to 1.05 for current reading |  | 1.0141
|  | 4 |  | 3.179 |  | 3.153 |  | 1.02 | ECL change to 1.06 for current reading |  | 1.0082
|  | 4.5 |  | 3.576 |  | 3.563 |  | 1.06 |  |  | 1.0036
|  | 5 |  | 3.975 |  | 3.971 |  | 1.14 |  |  | 1.001
|  | 5.5 | [6 empty] | overload voltage
Increase of Internal Resistance | 22.2 ohms | Set Point | Analog Reading | Monitor Reading | Analog Reading | Monitor Current | Corrected Mon. I | ECL | ECL | Notes
Multimeter voltage reading off by a factor of 10, more accurate to read off analog |  | [kV] | [kV] | [-kV] | [mA] | [V] | [mA] | [kV] | [mA]
Module 2 |  | 0 | 0 | 0 | 0 | 0 |  | 0.9 | 107 |  |  |  | Voltage | Current
Channel A Test 2 |  | 1 | 0.4 | 0.788 | 0.02 | 0.366 | 0.0366 | 0.9 | 108 | ECL change to 113 for current |  | 1 | 0.788 | 0.0366
|  | 2 | 1.1 | 1.582 | 0.05 | 0.743 | 0.0743 | 0.9 | 114 | ECL fluctuate between 114/115 for current |  | 5 | 3.972 | 0.1889
|  | 3 | 2 | 2.38 | 0.09 | 1.129 | 0.1129 | 0.9 | 114 | ECL fluctuate between 114/115 for current |  | 10 | 7.96 | 0.399
|  | 4 | 2.9 | 3.176 | 0.14 | 1.514 | 0.1514 | 0.9 | 115 | ECL fluctuate between 114/115 for current
|  | 5 | 3.8 | 3.972 | 0.18 | 1.889 | 0.1889 | 0.9 | 115
|  | 6 | 4.5 | 4.77 | 0.22 | 2.28 | 0.228 | 0.9 | 116
|  | 7 | 5.5 | 5.57 | 0.26 | 2.69 | 0.269 | 0.9 | 116 | ECL fluctuate between 116/17 for current
|  | 8 | 6.1 | 6.37 | 0.31 | 3.12 | 0.312 | 0.9 | 116 | ECL change to 117
|  | 9 | 7 | 7.16 | 0.35 | 3.55 | 0.355 | 0.9 | 117
|  | 10 | 7.9 | 7.96 | 0.39 | 3.99 | 0.399 | 0.9 | 117
V=IR
R=V/I
2023-12-14 |  | Set Point | Analog Reading | Monitor Reading | Analog Reading | Monitor Current | Corrected Mon. I | ECL | ECL | Notes
11.24am |  | [kV] | [kV] | [kV] | [mA] | [V] | [mA] | [Kv] | [mA]
Zubayda, Ian, Muhammad |  | 1 | 0.8 | 0.797 | 0.037 | 0.389 | 0.0389 | 0.88 | 0.88
Module34 (second), Channel B Test 1 |  | 5 | 3.8 | 3.982 | 0.19 | 1.933 | 0.1933 | 0.88 | 0.88
|  | 10 | 8 | 7.97 | 0.38 | 3.933 | 0.3933 | 0.87 | 0.88
Volatge and monitor threshold adjusted to max |  | [mA]: 0
Serial No. 8901282
Type. HT8176
2024-01-16
Zubayda, Lucia, Muhammag, Ian
5

## 33 Channel A&B
3.20pm | Potentiometer Reading | Analog V | Monitor V | Analog I | Monitor I | Calibrated M. I | Polarity | R |  | CMI | Mon. V | Overload Check
Channel 33 A | Points | kV | V | mA | V | V=mA |  | ohms | M ohms | A | Kv
Zubayda | 1 | 0.3 | 0.807 | 0.03 | 0.011 |  | negative
| 5 | [5 empty] | negative
| 8 | [5 empty] | negative
NOTES
Overload light on during ramp up from point 1 to 5. Reset and ramped up to point 2 for test, overload light on. Reset and ramped up to point 1.5, overload light on.
Monitor V is on module limit hence very high, does not equal Analog V.
Limits have put to max
Monitor I point 1 repeated, slow incline to 0.01 then overload light on, different to try 1, analog I still reads the same
| Potentiometer Reading | Analog V | Monitor V | Analog I | Monitor I | Calibrated M. I | Polarity | R |  | CMI | Mon. V | Overload Check
Channel 33 B | Points | kV | V | mA | V | V=mA |  | ohms | M ohms | A | Kv
4pm | 1 | 0.9 | 0.806 | 0.01 | 0.001
Zubayda | 2 | 1.8 | 1.621 | 0.05 | 0.008
3
NOTES
Overload light on during ramp up from point 1 to 5. Reset and ramped up to point 2 for test, overload light off, then ramped up to 3, overload light on
Limits have been put to max

## 34 Channel A&B
Channel 34 A
4.20PM | Potentiometer Reading | Analog V | Monitor V | Analog I | Monitor I | Calibrated M. I | Polarity | R
Zubayda | Points | kV | V | mA | V | V=mA |  | ohms | M ohms
| 1 | 0.5 | 0.789 | 0.03 | 0.428 |  |  |  |  | CMI | Mon. V | Overload Check
| 5 | 3.9 | 3.973 | 0.2 | 1.97 |  |  |  |  | A | Kv
| 10 | 8 | 7.93 | 0.38 | 3.97
NOTES
Point 5 Monitor I fluctuates regularly, does not settle ranged between 2.06-2.16. After 10 mins current decreasing

## 35 Channel A&B
2024-01-16
Zubayda, Muhammad, Lucia, Ian
5V fixed power supply in series with NIM crate, Current = 61mA
Switch to channel A from last year due to channel B unable to work
Protocol Implementation | Potentiometer Reading | Analog V | Monitor V | Analog I | Monitor I | Calibrated M. I | Polarity | R |  |  | CMI | Mon. V
3pm | Points | kV | V | mA | V | V=mA |  | ohms | M ohms |  | A | Kv
Channel 35 A | 1 | 0.8 | 0.796 | 0.03 | 0.36 | 0.036 | negative | 22111111.1111 | 22.1111 |  | 3.6e-05 | 796 | Overload Check
| 5 | 3.9 | 3.98 | 0.17 | 1.82 | 0.182 | negative | 21868131.8681 | 21.8681 |  | 0.000182 | 3980 | yes
| 8 | 8 | 7.99 | 0.36 | 3.64 | 0.364 | negative | 21950549.4505 | 21.9505 |  | 0.000364 | 7990
2.30PM | Potentiometer Reading | Analog V | Monitor V | Analog I | Monitor I | Calibrated M. 1 | Polarity | R |  |  | CMI | Mon. V | Overload Check
Channel 35 B | Points | Kv | V | mA | V |  |  | ohms | M ohms |  | A | Kv | yes
Zubayda & Mohammad | 1 | 0.6 | 0.787 | 0.03 | 0.37 | 0.037 | negative | 21270270.2703 | 21.2703 |  | 3.7e-05 | 787
| 5 | 3.9 | 3.973 | 0.16 | 1.88 | 0.188 |  | 21132978.7234 | 21.133 |  | 0.000188 | 3973
| 8 | 6.3 | 6.37 | 0.3 | 3.02 | 0.302 |  | 21092715.2318 | 21.0927 |  | 0.000302 | 6370
c

## 36 Channel A&B
19th Jan
Mohammad
note: Channel 36 A had a little tape on potentiometer
|  | Column1 | Column2 | Column3 | Column4 | Column5 | Column6 | Column7 | Column8 | Column9 | Column10 | Column11 | Column12
|  | Column1: 36 channel A | note: mon V and I are negative | overloaad check: | good
[5 empty] | Analog |  |  | Monitor |  |  | Monitor conversion
|  |  | potentiometer |  | Voltage | Current |  | Voltage | Current |  | Current | Voltage | R from graph
|  |  | point |  | kV | mA |  | V | mA |  | A | V | ohms
|  |  | 1 |  | 0.7 | 0.03 |  | 0.796 | 0.35 |  | 3.5e-05 | 796 | 21.9 Mohm
|  |  | 5 |  | 3.9 | 0.17 |  | 3.979 | 1.79 |  | 0.000179 | 3979
|  |  | 10 |  | 8 | 0.36 |  | 7.98 | 3.62 |  | 0.000362 | 7980
36 channel B
[5 empty] | Analog |  |  | Monitor
|  |  | potentiometer |  | Voltage | Current |  | Voltage | Current
|  |  | point |  | kV | A |  | V | A
|  |  | 1 |  | 1.5 | 0.07 |  | 1.56 | 0.69
|  |  | 5 |  | 4.8 | 0.22 |  | 4.75 | 2.13
|  |  | 10 |  | 8 | 0.36 |  | 7.98 | 3.6

## 37 Channel A&B
19th Jan
Mohammad
note: RAMP UP IS CONSIDERABLY FASTER FOR CHANNEL A&B
|  | Column1 | Column2 | Column3 | Column4 | Column5 | Column6 | Column7 | Column8 | Column9 | Column10 | Column11 | Column12
|  | Column1: 37 channel A | note: mon V and I are negative | overloaad check: | good
[5 empty] | Analog |  |  | Monitor |  |  | Monitor conversion
|  |  | potentiometer |  | Voltage | Current |  | Voltage | Current |  | Current | Voltage | R from graph
|  |  | point |  | kV | mA |  | V | mA |  | A | V | ohms
|  |  | 1 |  | 0.7 | 0.04 |  | 0.811 | 0.35 |  | 3.5e-05 | 811 | 21.9 Mohm
|  |  | 5 |  | 3.5 | 0.17 |  | 4 | 1.81 |  | 0.000181 | 4000
|  |  | 10 |  | 7.9 | 0.37 |  | 8 | 3.64 |  | 0.000364 | 8000
|  | Column1 | Column2 | Column3 | Column4 | Column5 | Column6 | Column7 | Column8 | Column9 | Column10 | Column11 | Column12
37 channel B
[5 empty] | Analog |  |  | Monitor |  |  | Monitor conversion
|  |  | potentiometer |  | Voltage | Current |  | Voltage | Current |  | Current | Voltage | R from graph
|  |  | point |  | kV | mA |  | V | mA |  | A | V | ohms
|  |  | 1 |  | 0.7 | 0.03 |  | 0.788 | 0.34 |  | 3.4e-05 | 788 | 21.9 Mohm
|  |  | 5 |  | 3.9 | 0.17 |  | 3.975 | 1.8 |  | 0.00018 | 3975
|  |  | 10 |  | 8 | 0.37 |  | 7.98 | 3.63 |  | 0.000363 | 7980

## 38 Channel A&B
19th Jan
Mohammad
note: noticably slow ramping?
|  | Column1 | Column2 | Column3 | Column4 | Column5 | Column6 | Column7 | Column8 | Column9 | Column10 | Column11 | Column12
|  | Column1: 38 channel A | note: mon V and I are negative | overloaad check: | good
[5 empty] | Analog |  |  | Monitor |  |  | Monitor conversion
|  |  | potentiometer |  | Voltage | Current |  | Voltage | Current |  | Current | Voltage | R from graph
|  |  | point |  | kV | mA |  | V | mA |  | A | V | ohms
|  |  | 1 |  | 0.8 | 0.03 |  | 0.8 | 0.35 |  | 3.5e-05 | 800 | 21.9 Mohm
|  |  | 5 |  | 3.9 | 0.16 |  | 4.02 | 1.82 |  | 0.000182 | 4020
|  |  | 10 |  | 8 | 0.36 |  | 8 | 3.65 |  | 0.000365 | 8000
5
|  | Column1 | Column2 | Column3 | Column4 | Column5 | Column6 | Column7 | Column8 | Column9 | Column10 | Column11 | Column12
38 channel B
[5 empty] | Analog |  |  | Monitor |  |  | Monitor conversion
|  |  | potentiometer |  | Voltage | Current |  | Voltage | Current |  | Current | Voltage | R from graph
|  |  | point |  | kV | mA |  | V | mA |  | A | V | ohms
|  |  | 1 |  | 0.5 | 0.03 |  | 0.794 | 0.35 |  | 3.5e-05 | 794 | 21.9 Mohm
|  |  | 5 |  | 3.9 | 1.17 |  | 3.984 | 1.8 |  | 0.00018 | 3984
|  |  | 10 |  | 8 | 0.37 |  | 8 | 3.64 |  | 0.000364 | 8000

## ramp speed analysis
check comment |  | capacitor on 271d | resistor
| 36 | 0.033uF | 15Mohms
| 37 | 0.1 uF
| 38 | 3.3 uF
Note on A1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Checking ramp speed by opening last 3 modules to see what the differences are