## Slide 1: Old and New Spacer Pop Tests
Calibration of 0-40mbar & 0-200mbar Pressure sensors.
Water column & Digital pressure gauge calibration.
Test to destruction of Old and New Kodel gaps.
Leak and “Popped Spacer Test”.
Conclusion.

## Slide 2
Chart: 
| -0.04 | 2.4500000000000002 | 6.32 | 14.99 | 17.45 | 24.05 | 29.32 | 32.17 | 36.479999999999997 | 40.549999999999997 | 30.44 | 24.33 | 9.07 | 0.16500000000000001
 | 0 | 2.6 | 6.4 | 14.2 | 17.8 | 24.2 | 29.5 | 32.4 | 36.6 | 40.6 | 30.6 | 24.3 | 9 | 0
24 bit ADC
Pico Recorder
Sensor Technics
CTEM 7000 and9000
Calibration of sensors & ADC against WC

## Slide 3
Method and plots of destructive Tests
Initial failure , the first spacer
2000
1000
Time [s]
Pressure [mbar]
100
50
0
Remaining spacer failures
An example
RE4/2 TW01

## Slide 4
Individual spacer ID
Evidence of a series of 5
5 ramps and failures.
RE3/2 TW18
Propagation of bands of 5 spacers across gap width

## Slide 5
Results
gap name | test date | Initial Failure pressure |  | Comments |  | 
 |  | [mbar] |  |  |  | 
RE3/3 - TW19 | 10-Dec-2010 | 100 |  | Digital P. gauge. |  | 
RE2/1 - TW16 | 10-Dec-2010 | 100 |  | Digital P. gauge |  | 
RE4/3 - TW02 | 10-Dec-2010 | 16.6 |  | WC. ~24 spacers "gone" before starting |  | 
RE4/2 – B02 | 14-Jan-2011 | 94 |  | ADC. 'spacer test" |  | 
RE4/2 – TW01 | 13-Jan-2011 | 120 |  | ADC |  | 
RE1/2 – B12 | 11-Jan-2011 | 116 |  | ADC, error < 40micro bar |  | 
RE1/3 – TN12 | 12-Jan-2011 | 153 |  | ADC, error < 40micro bar |  | 
RE3/2 – TW18 | 11-Jan-2011 | 146 |  | ADC |  | 
mbar
Chart: 
| •RE3/3 - TW19 | •RE2/1 - TW16 | •RE4/3 - TW02 | •RE4/2 – B02 | •RE4/2 – TW01 | •RE1/2 – B12 | •RE1/3 – TN12 | •RE3/2 – TW18
Pressure to Initial Failure | 100 | 100 | 16.6 | 94 | 120 | 116 | 153 | 146
Operating pressures are 2-3 mbar . QA tests are done @ 20mbar, which will not prejudice the integrity of the bond
Numerous spacers already broken

## Slide 6
“Spacer test” and leak test
done using ADC
“Standard finger thrust” gives approx 0.1-0.2mbar. Equivalent to 1-2mm WC
Slope gives leak rate
Spacers can be counted to check for errors in procedure
Unglued or popped spacers will give higher values

## Slide 7: Conclusions
Two of the 3 agreed “new” gaps were sacrificed in addition
to “old” gaps stored in the ISR, to establish pressure to failure.
The bonding of spacers in the “new” production from KODEL appear sufficient for our needs and similar to previous production.
The use of an online log of pressure provides a reliable and
simple popped spacer test combined with a leak test.
The reason for two “new” gaps having popped spacers at
less than 20 mbar has not been clarified.
Thanks to Abeera and Luigi