                                             Gap production  specification  document
1 Introduction

A total of 660 Gas Volumes (GV) are needed for the RPC low η restoration. For continuity with the existing CMS forward RPC system,  the production will occur at the KODEL laboratory (Korea University, Seul) using the same technology as developed in past for the CMS production. The HPL panels needed for the gap assembly will be procured in Europe, cut to the proper size,  surface cleaned and shipped to KODEL. Procurement of other basic components, as described in section 2,  will be pursued directly by KODEL. 
In the following table 1 the list of gap typologies  and associated  dimensions  is given. 
                                     Table 1: list of gap types and dimensions 
	Gap Name
	Quantity
	Dimension
	URL engineer drawing 

	RE4/3 “top narrow”
	110
	
	

	RE4/3 “top wide”
	110
	
	

	RE4/3 “bottom”
	110
	
	

	RE4/2 “top narrow”
	110
	
	

	RE4/2 “top wide”
	110
	
	

	RE4/2 “bottom”
	110
	
	


In annex 1 of the present document  engineering  drawings  for  each gap type gives  construction details  concerning  spacers position and  electrical and gas connections.  

2 Procurement and preparation of  basic components  
2.1  List of components

The following  basic components should be procured  under responsibility of KODEL
· Graphite paint

Name. DAG 305 solved in MIBK
Produced by: Akenson
Drying time: 1 hour

· Insulating films

Material specification: Polyetilenterephatale

Name: DIEPOL
Produced by: Effegidi S.P.A.

Thickness: 190 m

Dielectric rigidity ( @25 0C-50 Hz) 150 KV/mm (ASTM D149)

Dielectric constant ( @25 0C-50 Hz) 3.3 (ASTM D150)

Volume resistivity ( @25 0C-DC) 1018 (ASTM D257)

Surface resistivity ( @25 0C-50 Hz) 1014 (ASTM D257)

Isolation class: B (DIN 57 530)

Melting temperature: 260 0C
· Hot melt adhesive

Name: TERMOMELT PS 856
Produced by: LEUENBERGER & c S.P.A.

Working temperature: 150-170 0C

· Spacer edge frame and buttons:

Material name: Polycarbonate

Production method: Injection molding

Thickness and button: 2.000+/-0.015 mm
· Gas inlet/outlet:

Material name: Polycarbonate

Production method: Injection molding

Thickness: 2.000+/-0.015 mm

· Epoxy adhesive 

Name: Scotch Weld DP 460
Produced by: 3M
Work life (@ 23 0C): 60-90 min

Min. polymerization time (@ 23 0C):4-6 hours

Max shear stress with polycarbonate: 2.8-3.4 Mpa (after 24 hours)

10% (after 5 hours)

25% (after 6 hours)

90% (after 7 hours)
· Linseed Oil

……………

…………….
· Insulating  HV tape
……………………
……………………..
· HV cable
……………………

……………………..
2.2  Components certification and preparation

………………………………………….
3 Production  procedure
3.1  Preparation of electrodes
The graphite paint is diluted in Methyl-Isobutyl-Khetone with the proportion XX/YY. The HPL plates to be painted are masked with paper tape along the edges to cover a  width of 9 mm. The graphite mixture is applied by means of  xerography  method. At the end of the procedure the plate is kept XXX minutes in an air stream until the solvent has fully evaporated. The graphite layer is thickened, along the plate with a further coat of paint applied with a brush. The masking tape finally is removed. The resistivity of each plate is measured between the two middle positions of the short  sides of the trapezoid. The resistance values must be within the value range reported for each gap  type  in table 2.

              Table 2 : ranges of allowed resistance values

	Gap Name
	Rmin (K
	Rmax(K

	RE4/3 “top narrow”
	80
	150

	RE4/3 “top wide”
	80
	150

	RE4/3 “bottom”
	80
	150

	RE4/2 “top narrow”
	80
	150

	RE4/2 “top wide”
	80
	150

	RE4/2 “bottom”
	80
	150


The coated electrode is isolated from the external readout panels with a PET film, 190 m thick, which is glued on the graphite with a “hot melt” type adhesive. The adhesive is heated above the melting point and then pumped from the melting point up to an extrusion head which is kept at constant, well-controlled temperature. The PET film is rolled at constant velocity by a rubber coated rotating iron cylinder, which keeps it in contact with the extrusion head so that an uniform layer of hot adhesive (50-70 mm thickness) is deposited on its surface. 

One barcode on each electrode  plus should be put on  the PET film placed  near to the edge but away from the foreseen  electrical and gas connections.
3.2  Gas volume assembly
Two electrodes, prepared as described in previous step,  are glued together using polycarbonate spacer and edge frames to form a gas volume (GV). An epoxy adhesive is used. Information about pattern of the spacers, edge frame and gas inlet/outlet components are available in Annex 1.
At each step, the following operation are performed in parallel for all spacers: 1)deposition of drops of adhesive, 2)deposition of spacers, 3)deposition of a second drops of adhesive on top of the deposed spacers. The edge frame is glued with the same procedure. A thin bead of glue is applied all along the edge. The frame segments are put in position, one in each side of the plate, and a second string of glue is put on top of the frame. The second panel is then overlapped to the first one.

During the curing time the gas volume is kept under a constant pressure  distributed uniformly along the entire surface  of  ZZZ bar at constant temperature for  24 hours.
Before the gluing a careful  examination of  HPL surface should be performed to select only appropriate  panels and degreasing procedure should be applied, if needed, of all pieces. The gluing must take place at room temperature with relative humidity below the 50%. Excessive level of humidity could end up in poor glue curing.
The two components of the adhesive are mixed together only inside the dispensers, i.e. very short time before the application. Relevant parameters for this step are:  1)glue workable time after components are mixing should not exceed 60 min; 2)typical time between components mixing and pressure application should less than 20 min.
Every ten GV a spacer is glued on a small test plate and cured during the 8 hours under standard conditions. After that time, the spacer is gradually stressed for 10 min under controlled conditions up the breaking point. The test is successful if the breaking occurs with a normal stressing force of over 8 Kg. In case of failure of the gluing, the last ten GV assembled before test are discharged.
Final sealing of the GV is done filling with hot melt the 2x3 mm section that remain around the edges of the frame. The excess of melt is cut out after it has cooled down to room temperature. 
3.3  Gas Volumes varnishing
A batch of 20 GV, sandwich between two rigid panels, is placed in vertical position with the long side inclined at about 100 mrad with respect the horizontal plane. The gas inlets are connected with a plastic flexible pipes to a steel cylinder containing the which can be moved along the vertical direction.  Varnish composition should be linseed oil/ hetane (40/60).  The surface of the liquid inside the cylinder lies initially below the lower GV inlets. The cylinder is then lifted at constant speed so the varnish level visible through the transparent pipes. The varnish circuit is sealed and, during the filling, the pressure inside the GV is kept negative with respect to the atmosphere to compensate the varnish static pressure. The cylinder is subsequently lowered at very low speed up to the complete emptying of the gas volume. Depletion rate in this phase should be  3 cm/min.  At the end, the GV are connected to airflow (30-40 volume/hour, Tair = XXX)  for 36 hours up to complete polymerization of the varnish.
Every varnishing cycle, a single GV is open and inspected to ensure the quality of the procedure.  GV internal surface should be checked for complete polymerization  of the oil coating,  for absence of  anomalous oil drops around the spacer and for other imperfections.  Moreover the thickness  of the oil coating should be evaluated not to be more then 30 +/- 10 m. In case one of the above mentioned criteria is not met, the full batch is discharged.
T the end of this step, if successful, one barcode label specifying the GV  ID  should be placed  near to the edge of each gap but  away from foreseen electrical and gas connections.
4 Quality assurance

Every single gap is tested for gas and spacers tightness and for high voltage. Spacers and gas tightness are performed with Argon45.  Electrical tests are performed with the standard CMS mixture  of C2H2F4 (95%), iso-C4H10 (5 %). The gas mixture is humidified at the level of 40% (??).
4.1 “Gas leakage” test

The gas tightness is controlled by measuring the stability of an applied overpressure of 20mb  in a fixed time of 200 seconds using a pressure sensor reference xxx and a ADC pico loger recorder ref. xxxxx of 20 bit resolution. According to the gap type, the following upper leak rates are considered:
	Gap Type
	RE4/3 “top narrow”
	RE4/3 “top wide”
	RE4/3 “bottom”
	RE4/2 “top narrow”
	RE4/2 “top wide”
	RE4/2 “bottom”

	Maximum Leak Rate
	
	
	
	
	
	


An additional measurement of the leak rate has to be curry out at 3 mbar, to be used as a reference value for the final commissioning  phase at P5. 

4.2  “Spacer strength” test 

The GV are subjected to an overpressure of 2 0mbar in order to check the spacers tightness. …………………

GV with at least one spacer detached are rejected.
4.3  “I vs HV characterization”

All the GV passing this test  will then be eligible  to be connected to the HV power supply for further certification. To this aim two additional steps are needed:

· the full perimeter of the GV should be protected against electrical leakage using proper HV insulating  tape;

· HV connections should be performed (using proper cable)  on the small copper pads inserted below the  insulating PET film.  The soldering should be protected using  hot melt  and insulating  HV tape.

The electrical cables should be  diameter  3 mm, color red, 20 KV cable for the High Voltage;  0.5 mm2 conductor cross section, color black,  800 V cable for the ground. Both cables will have a length of 1.5m. A Jupiter connection should be connected  using heat shrink (diameter of 10 mm) and a shield (diameter xxx). The central electrode should be connected to the red (HV) cable.  The black, 0V line and the shield soldered to the Jupiter casing.

Electrical tests have to be done with the CAEN module A1526N using one channel per GV.
“Dark current” test (I) 
Measure the current at different HV values and plot the CURRENT vs  HVeff   (HVeff  = 0.1, 1.0, 2.0, 3.0, 4.0, 5.0, 6.0, 7.0, 7.5, 8.0, 8.5, 9.0, 9.1, 9.2 , 9.3, 9.4, 9.5, 9.6, 9.7, 9.8, 9.9, 10.0, 10.1 KV, 23 points in total). Each voltage step shall be applied for duration of  30 minutes  before the measurement of the current. Pressure, T and Humidity have to be monitored and  of HVeff  should be calculated according the usual P,T correction factor (XXXX) to normalize values  at  1010 mbar pressure  and 293 °K temperature.
If  the current at HVeff  =10 kV is higher than  Imax  the  GV is rejected.   According to the GV type  the following Imax are considered:
	Gap Type
	RE4/3 “top narrow”
	RE4/3 “top wide”
	RE4/3 “bottom”
	RE4/2 “top narrow”
	RE4/2 “top wide”
	RE4/2 “bottom”

	Imax
	
	
	
	
	
	


“Dark current”  test (II) 

Leave the GV under HVeff =  10 kV for a period of 168 hours with a gas flow of 5 l/hour . Monitor the pressure and temperature,   correct the applied voltage online with a cycle of 5 minutes in order to maintain the HVeff constant  at 10000V.  The gap will be rejected if at the end of this period anyone of the following criteria have been met: I10KV  > Imax,  current increase  higher than  50% , an OVC  or a trip occurred. 

“Dark current”  test (III) 
Measure the current at the following  HVeff values (HVeff =1.0, 2.0, 3.0, 4.0, 4.5, 5.0, 5.5, 6.0, 6.5, 7.0 kV, 10 points in total) , and plot the HVeff vs Current. Each step is  of 5 minutes duration. Calculate the slop which give you the resistance and the fitted value at 6 KV. If the level of the current at 6 KV effective is less than 0.4 uA (????) the gap is rejected If the slop is less than …???? … the gaps is rejected

5 Storage, shipment  and certification

The builder needs to have an efficient system for storage and tracking of the raw materials and the finished parts, in order to avoid damage and contamination of the surface in the parts and mixed-up in the raw materials. It is important to avoid that the HPL  plates and GV are subject to a bending moment due to the gravity, this occurs maintaining the structures horizontally. This must be avoided especially for a long period of time, because can procure a permanent deformation on the bakelite plate, modules and chamber.

Storage should be such that the gaps will never be stressed in bending. Humidity and Temperature of the storage area should be recorded  and  kept between 18  and 25 degrees Celsius with a  relative Humidity between 35% and 50 % .

A final visual inspection of the gaps before their packing should  be done.

6 Schedule

                  ANNEX 1a

                  Gap dimension and engineering details  for RE4/2
[image: image1.emf]
[image: image2.emf]
    ANNEX 2a

                  Gap dimensions and engineering details  for RE4/3
[image: image3.emf]
